## Slide 1: PFNA(Proximal Femoral Nail, Anti-rotation)

## Slide 2: Anatomy of Proximal Femur
Subtrochanteric
Intertrochanteric
Femur neck
Common types of hip fractures

## Slide 3: Fractures types
Stable
Unstable
A1:simple trochanteric fractures:
A2 multifragmentary trochanteric fractures:
A3 intertrochanteric fractures:
Lost of Medial Buttress
No lateral Wall
For Stable fractures, Plates & Nail both can be used
For Unstable fractures, it is recommended to use IM nail

## Slide 4: Advantage of IM nail
Bear less loading force
  Motion of Force = F* D/d
  D > d
Mini Invasive
IM nail is the first choice for the unstable fractures

## Slide 5: Features of PFNA
Long
Standard
Several distal locking options
PFNA Lag screw
  Compaction of cancellous bone, especially for elderly & osteoporosis patients
  Higher cut-out resistance
  Anti-rotation
5°
Optimal fit: medial-lateral angle of 5°
This allows insertion at the tip of the greater trochanter
Notes: Static locking: usually used for comminuted fractures,
Dynamic locking: usually used for transverse or distracted fractures

## Slide 6: Components of PFNA
short
Long
End Cap
PFNA Lag screw
Femoral Nail Locking Screw
Nail

## Slide 7: Indications
Short
Long
PFNA short (Length 170 mm –240 mm)
Pertrochanteric fractures
Intertrochanteric fractures
High subtrochanteric fractures
PFNA long (Length 300 mm –420 mm)
Low and extended subtrochanteric fractures
Ipsilateral trochanteric fractures
Combination fractures (in the proximal femur)

## Slide 8: Patient Position & Fracture Reduction
Place the patient on the supine position.
.Adduct the affected leg by 10 –15°
Perform closed reduction of the fracture under image intensifier control. Perform open reduction when necessary

## Slide 9: Confirm nail length and diameter
Use the Radiographic Ruler to determine the length and the diameter of the nail under the C-arm

## Slide 10: Approach
Make a 5 cm incision proximal from the tip of the greater trochanter

## Slide 11: Open Femur
AP view
Step 1 Entry Point:
  AP view: on the tip or slightly lateral to the tip of the greater trochanter
  Lateral view: in line with the axis of the intramedullary canal
Step 2 Insert Guide Wire
  Insert the guide wire through the protection sleeve and the drill sleeve
Step 3 Open Femur with awl

## Slide 12: Open Femur(continue)
Step 4 Open Femur
  Guide the cannulated drill bit through the protection sleeve over the guide wire and drill to the stop on the protection sleeve
Step 5 Ream medullary canal
  Ream the cannal with the flexible reamer and the reamer head
  It is recommended to start with the 8.5 mm diameter reaming head, ream to a diameter of 0.5 to 1.5 mm greater than the nail diameter

## Slide 13: Insert Nail
Step 1 Assemble PFNA
Step 2 Insert PFNA
  Insert the PFNA with the Hammer Guide and the Slotted Hammer

## Slide 14: Proximal Locking
Step 1 Prepare Guide Wire insertion:
  Fix the aiming arm firmly to the insertion handle
  Insert the sleeve assembly to the aiming arm
Step 2 Insert Guide Wire
  Insert a new guide wire through the drill sleeve into the bone
  The optimal position of the guide wire in the 1/3 part of the middle low femoral neck and at a distance of 5-10 mm below the joint level
Step 3 Measure the length
  The measuring device indicates the exact length of the guide wire in the bone
Notes: The sleeve assembly must be in contact with the bone during the entire blade implantation
Insert the guide wire for the PFNA blade carefully to avoid penetration into the joint

## Slide 15: Proximal Locking(Continue)
Step 4 Open lateral cortex
  Push the cannulated drill bit Ø3.2 over the guide wire. Drill to the stop
Step 5 Drill for PFNA blade
  Set the chosen blade length of the measure length by fixing the fixation sleeve in the corresponding position
  Push the reamer over the guide wire. Drill to the stop

## Slide 16: Insert PFNA Lag Screw
Step 6 Assemble PFNA blade
  Insert the PFNA blade into the key and then attach them to the impactor. Screw the impactor counterclockwise to unlock
Step 7 Insert PFNA blade
  Insert the PFNA blade to the stop by applying gentle blows with the hammer
Step 8 Lock PFNA blade
  To lock the PFNA blade, turn the impactor clockwise and tighten the blade
Unlocked
Locked
Notes: The PFNA blade is supplied in a locked state.

## Slide 17: Distal Locking for PFNA(short)
Step 1 Assemble distal aiming arm
  Static locking and dynamic locking are available for standard PFNA
  Fix the chosen aiming arm firmly to the insertion handle
Step 2 Drill & Measure the depth
  Use the drill bit to drill through both cortices
  Measure the depth with the Depth Gauge
Step 3 Insert screw
  Insert a femoral locking screw of the measured length with the screwdriver

## Slide 18: Distal Locking for PFNA(long)
Step 1 Assemble distal aiming arm
  Assemble the distal aiming arm according to the selected nail length & L/R
  Attach the distal target device to the distal aiming arm
Step 2 calibrating
  Drill with the drill bit Ø5.2 until to the PFNA platform
  Insert the Positioning Rod and fix it to the platform with the Locking Block
Step 3 Insert locking screw
  Insert a femoral locking screw of the measured length with the screwdriver
  Static locking and dynamic locking can be chosen accordingly

## Slide 19: Insert End cap
Step 1 Remove instruments
  Remove all the connecting instruments
Step 2 Insert End Cap
  Screw the end cap into the proximal end of the nail and tighten it firmly

## Slide 20: Implant Removal
Step 1 Remove end cap
  Remove the end cap with the Hexagonal screwdriver
Step 2 Remove PFNA Lag Screw
  Extract the PFNA Lag Screw with the Extraction Screw
Step 3 Remove distal locking screw and nail
  Remove the distal locking screw with the Hexagonal screwdriver
  Screw the hammer guide into the PFNA and tighten it and Extract the nail by applying gentle blows with the slottled hammer

## Slide 21: Specification
PFNA(short)
Diameter: φ9.5, φ10, φ11&φ12
Length: 170mm, 200mm, 240mm
PFNA(long)
Diameter: φ9.5, φ10, φ11
Length: 300-420 mm Increased every 20mm, L/R
PFNA Lag Screw
Diameter: φ10.5
Length: 75-120 mm Increased every 5mm
Femoral Nail Locking Screw
Diameter: φ4.5
Length:25-80mm, Increased every 5mm
PFNA End Cap
Length: 0, 5, 10, 15

## Slide 22
Thank You